ROF ABC-MZ

Automatic bias point control module

ROF-ABC-MZ series automatic bias control module is used for automatic bias control of lithium niobate M-Z modulator, which can make the modulator work stably at the lowest point, the highest point or the orthogonal point (linear region). The module is also integrated with 1/99 coupler, which can control the switching of the working point through the external serial port, and supports manual adjustment mode, which is suitable for a variety of different wavelength modulators and applications, and is very suitable for university laboratories to build desktop experiments.

Feature

Application

- Multiple bias operating modes are available (Quad+↔Quad-, Min ↔Max)

- Serial communication, programmed automatic fine tuning and locking bias points

- Internal component beamers support a variety of wavelengths

- Module package, adapter power supply

Application

- Optical fiber communication

- Microwave photon

- Pulsed light application

Parameters

| Argument | Min | Typical | Max | Unit |
| --- | --- | --- | --- | --- |
| Optical parameter |  |  |  |  |
| Input optical power 1* | 0 |  | 13 | dBm |
| Operating wavelength 2* | 780 |  | 1650 | nm |
| Optical fiber interface | FC/APC |  |  |  |
| Electrical parameter |  |  |  |  |
| Bias voltage | -10 |  | 10 | V |
| Switch extinction ratio 3* | 20 | 25 | 50 | dB |
| Mode-locked region | Positive or Negative |  |  |  |
| Lock mode | Quad+ (Quad-) or Min (Max) |  |  |  |
| Modulation depth (QUAD) |  | 1 | 2 | % |
| Modulation depth (Null) |  |  | 0.1 | % |
| Pilot frequency (QUAD) |  | 1K |  | Hz |
| Pilot frequency (NULL) |  | 2K |  | Hz |
| Conventional parameter |  |  |  |  |
| Dimensions (length × width × thickness) | 120×70×34 mm |  |  |  |
| Operating temperature | 0 - 70℃ |  |  |  |

Note:

1* refers to the power range input to the module when the modulator output is maximum. For the low point control of the modulator with a high extinction ratio, the input power should be appropriately increased; With special power input requirements, you can adjust the internal coupler and detector gain indicators, please consult sales when placing orders.

2* When placing an order, please specify the working wavelength, which needs to be optimized according to the working wavelength.

3* switching extinction ratio also depends on the switching extinction ratio level of the modulator itself.

Size drawing (mm)

Ordering information

| ROF | ABC | Modulator type | XX | XX | XX |
| --- | --- | --- | --- | --- | --- |
|  | Automatic bias point control module | MZ---M-Z modulator | Working wavelength: 15---1550nm 13---1310nm 10---1064nm 08---850nm 07---780nm | Fiber type: S-- Single mode optical fiber P - polarization-maintaining fiber | Optical fiber interface: FA—FC/APC FP---FC/UPC |

*If you have any special requirements, please contact our sales staff